## Slide 1: Engineering 80 – Spring 2015 Temperature Measurements
SOURCE: http://elcodis.com/photos/19/51/195143/to-92-3_standardbody__to-226_straightlead.jpg
SOURCE: http://www.accuglassproducts.com/product.php?productid=17523
SOURCE: http://www.eng.hmc.edu/NewE80/PDFs/VIshayThermDataSheet.pdf

## Slide 2: Key Concepts
  Measuring Temperature
  Types of Temperature Sensors
    Thermistor
    Integrated Silicon Linear Sensor
    Thermocouple
    Resistive Temperature Detector (RTD)
  Choosing a Temperature Sensor
  Calibrating Temperature Sensors
  Thermal System Transient Response

## Slide 3: What is Temperature?
SOURCE: http://www.clker.com/cliparts/6/5/b/f/11949864691020941855smiley114.svg.med.png

## Slide 4: What is Temperature?
AN OVERLY SIMPLIFIED DESCRIPTION OF TEMPERATURE
"Temperature is a measure of the tendency of an object to spontaneously give up energy to its surroundings. When two objects are in thermal contact, the one that tends to spontaneously lose energy is at the higher temperature.“
(Schroeder, Daniel V. An Introduction to Thermal Physics, 1st Edition (Ch, 1). Addison-Wesley.)
units
SOURCE: http://hyperphysics.phy-astr.gsu.edu/hbase/thermo/temper2.html#c1

## Slide 5: What is Temperature?
A SIMPLIFIED DESCRIPTION OF TEMPERATURE
"Temperature is a measure of the tendency of an object to spontaneously give up energy to its surroundings. When two objects are in thermal contact, the one that tends to spontaneously lose energy is at the higher temperature.“
(Schroeder, Daniel V. An Introduction to Thermal Physics, 1st Edition (Ch, 1). Addison-Wesley.)
units
SOURCE: http://hyperphysics.phy-astr.gsu.edu/hbase/thermo/temper2.html#c1

## Slide 6: Measuring Temperature with Rockets

## Slide 7: Measuring Temperature with Rockets
What are desirable characteristics of a temperature sensor?

## Slide 8: Desirable Temperature Sensor Characteristics
ACCURATE
REPEATABLE
WIDE TEMPERATURE RANGE
EASY CALIBRATION
FAST RESPONSE
SIMPLE RELATIONSHIP
SENSOR OUTPUT  TEMPERATURE
TEMPERATURE SENSOR
COST

## Slide 9: Thermistor
Thermistor – a resistor whose resistance changes with temperature

## Slide 10: Thermistor
Thermistor – a resistor whose resistance changes with temperature
Resistive element is generally a metal-oxide ceramic containing Mn, Co, Cu, or Ni
Packaged in a thermally conductive glass bead or disk with two metal leads

## Slide 11: Thermistor
Thermistor – a resistor whose resistance changes with temperature
Resistive element is generally a metal-oxide ceramic containing Mn, Co, Cu, or Ni
Packaged in a thermally conductive glass bead or disk with two metal leads
Suppose we have a “1 kΩ thermistor”…
  What does this mean?

## Slide 12: Thermistor
Thermistor – a resistor whose resistance changes with temperature
Resistive element is generally a metal-oxide ceramic containing Mn, Co, Cu, or Ni
Packaged in a thermally conductive glass bead or disk with two metal leads
Suppose we have a “1 kΩ thermistor” …
  What does this mean?
  At room temperature, the resistance of the thermistor is 1 kΩ

## Slide 13: Thermistor
Thermistor – a resistor whose resistance changes with temperature
Resistive element is generally a metal-oxide ceramic containing Mn, Co, Cu, or Ni
Packaged in a thermally conductive glass bead or disk with two metal leads
Suppose we have a “1 kΩ thermistor”
  What does this mean?
  At room temperature, the resistance of the thermistor is 1 kΩ
  What happens to resistance as we increase temperature?

## Slide 14: Negative Temperature Coefficient
Most materials exhibit a negative temperature coefficient (NTC)
  Resistance drops with temperature!

## Slide 15: Converting Resistance to Temperature
The Steinhart-Hart Equation relates temperature to resistance
SOURCE: http://p.globalsources.com/IMAGES/PDT/B1055847338/Thermistor.jpg

## Slide 16: Converting Resistance to Temperature
The Steinhart-Hart Equation relates temperature to resistance
T is the temperature (in Kelvin)
R is the resistance at T and Rref is resistance at Tref
A1, B1, C1, and D1 are the Steinhart-Hart Coefficients
HOW COULD WE DETERMINE THESE COEFFICIENTS?
SOURCE: http://p.globalsources.com/IMAGES/PDT/B1055847338/Thermistor.jpg

## Slide 17: Converting Resistance to Temperature
The Steinhart-Hart Equation relates temperature to resistance
T is the temperature (in Kelvin)
R is the resistance at T and Rref is resistance at Tref
A1, B1, C1, and D1 are the Steinhart-Hart Coefficients
HOW COULD WE DETERMINE THESE COEFFICIENTS?
  Take a look at the data sheet
SOURCE: http://p.globalsources.com/IMAGES/PDT/B1055847338/Thermistor.jpg

## Slide 18: Converting Resistance to Temperature

## Slide 19: Converting Resistance to Temperature

## Slide 20: Converting Resistance to Temperature
WHAT IF YOU LOST THE DATA SHEET, DON’T BELIEVE IT, OR WOULD LIKE TO VERIFY THE VALUES?

## Slide 21: Converting Resistance to Temperature
The Steinhart-Hart Equation relates temperature to resistance
T is the temperature (in Kelvin)
R is the resistance at T and Rref is resistance at Tref
A1, B1, C1, and D1 are the Steinhart-Hart Coefficients
HOW COULD WE DETERMINE THESE COEFFICIENTS?
  Take a look at the data sheet
SOURCE: http://p.globalsources.com/IMAGES/PDT/B1055847338/Thermistor.jpg

## Slide 22: Converting Resistance to Temperature
The Steinhart-Hart Equation relates temperature to resistance
T is the temperature (in Kelvin)
R is the resistance at T and Rref is resistance at Tref
A1, B1, C1, and D1 are the Steinhart-Hart Coefficients
HOW COULD WE DETERMINE THESE COEFFICIENTS?
  Take a look at the data sheet
  Measure 3 resistances at 3 temperatures
    Matrix Inversion (Linear Algebra)
SOURCE: http://p.globalsources.com/IMAGES/PDT/B1055847338/Thermistor.jpg

## Slide 23: Converting Resistance to Temperature
The Steinhart-Hart Equation relates temperature to resistance
T is the temperature (in Kelvin)
R is the resistance at T and Rref is resistance at Tref
A1, B1, C1, and D1 are the Steinhart-Hart Coefficients
HOW COULD WE DETERMINE THESE COEFFICIENTS?
  Take a look at the data sheet
  Measure 3 resistances at 3 temperatures
    Matrix Inversion (Linear Algebra)
  Least Squares Fit
SOURCE: http://p.globalsources.com/IMAGES/PDT/B1055847338/Thermistor.jpg

## Slide 24: How is Resistance Measured?
SOURCE: http://p.globalsources.com/IMAGES/PDT/B1055847338/Thermistor.jpg

## Slide 25: Thermistor Resistance (RT)
A thermistor produces a resistance (RT), which must be converted to a voltage signal

## Slide 26: Power Dissipation in Thermistors
A current must pass through the thermistor to measure the voltage and calculate the resistance
The current flowing through the thermistor generates heat because the thermistor dissipates electrical power
P = I2RT
The heat generated causes a temperature rise in the thermistor
This is called Self-Heating
WHY IS SELF-HEATING BAD?
I

## Slide 27: Power Dissipation and Self-Heating
Self-Heating can introduce an error into the measurement
The increase in device temperature (ΔT) is related to the power dissipated (P) and the power dissipation factor (δ)
P = δ ΔT
Where P is in [W], ΔT is the rise in temperature in [oC]
Suppose I = 5 mA, RT = 4 kΩ, and δ = 0.067 W/oC, what is ΔT?

## Slide 28: Power Dissipation and Self-Heating
Self-Heating can introduce an error into the measurement
The increase in device temperature (ΔT) is related to the power dissipated (P) and the power dissipation factor (δ)
P = δ ΔT
Where P is in [W], ΔT is the rise in temperature in [oC]
Suppose I = 5 mA, RT = 4 kΩ, and δ = 0.067 W/oC, what is ΔT?
(0.005 A)2(4000 Ω) = (0.067 W/oC) ΔT
ΔT = 1.5 oC
What effect does a ΔT of 1.5 oC have on your thermistor measurements?

## Slide 29: Power Dissipation and Self-Heating
Self-Heating can introduce an error into the measurement
The increase in device temperature (ΔT) is related to the power dissipated (P) and the power dissipation factor (δ)
P = δ ΔT
Where P is in [W], ΔT is the rise in temperature in [oC]
Suppose I = 5 mA, RT = 4 kΩ, and δ = 0.067 W/oC, what is ΔT?
(0.005 A)2(4000 Ω) = (0.067 W/oC) ΔT
ΔT = 1.5 oC
What effect does a ΔT of 1.5 oC have on your thermistor measurements?
How can we reduce the effects of self-heating?

## Slide 30: Power Dissipation and Self-Heating
Self-Heating can introduce an error into the measurement
The increase in device temperature (ΔT) is related to the power dissipated (P) and the power dissipation factor (δ)
P = δ ΔT
Where P is in [W], ΔT is the rise in temperature in [oC]
Suppose I = 5 mA, RT = 4 kΩ, and δ = 0.067 W/oC, what is ΔT?
(0.005 A)2(4000 Ω) = (0.067 W/oC) ΔT
ΔT = 1.5 oC
What effect does a ΔT of 1.5 oC have on your thermistor measurements?
How can we reduce the effects of self-heating?
  Increase the resistance of the thermistor!

## Slide 31: Thermistor Signal Conditioning Circuit
A voltage divider and a unity gain buffer are required to measure temperature in the lab

## Slide 32: Integrated Silicon Linear Sensors
An integrated silicon linear sensor is a three-terminal device
  Power and ground inputs
  Relatively simple to use and cheap
  Circuitry inside does linearization and signal conditioning
  Produces an output voltage linearly dependent on temperature
3.1 – 5.5 V

## Slide 33: Integrated Silicon Linear Sensors
An integrated silicon linear sensor is a three-terminal device
  Power and ground inputs
  Relatively simple to use and cheap
  Circuitry inside does linearization and signal conditioning
  Produces an output voltage linearly dependent on temperature
  When compared to other temperature measurement devices, these sensors are less accurate, operate over a narrower temperature range, and are less responsive
3.1 – 5.5 V

## Slide 34: Summary Thus Far…
>

## Slide 35: Thermocouple
Thermocouple – a two-terminal element consisting of two dissimilar metal wires joined at the end
SOURCE: http://upload.wikimedia.org/wikipedia/en/e/ed/Thermocouple_(work_diagram)_LMB.png

## Slide 36: The Seebeck Effect
Seebeck Effect – A conductor generates a voltage when it is subjected to a temperature gradient

## Slide 37: The Seebeck Effect
Seebeck Effect – A conductor generates a voltage when it is subjected to a temperature gradient
  Measuring this voltage requires the use of a second conductor material
Will I observe a difference in voltage at the ends of two wires composed of the same material?
Nickel-Chromium Alloy
Nickel-Chromium Alloy

## Slide 38: The Seebeck Effect
Seebeck Effect – A conductor generates a voltage when it is subjected to a temperature gradient
  Measuring this voltage requires the use of a second conductor material
  The other material needs to be composed of a different material
Nickel-Chromium Alloy
Copper-Nickel Alloy
The relationship between temperature difference and voltage varies with materials

## Slide 39: The Seebeck Effect
Seebeck Effect – A conductor generates a voltage when it is subjected to a temperature gradient
  Measuring this voltage requires the use of a second conductor material
  The other material needs to be composed of a different material
The voltage difference of the two dissimilar metals can be measured and related to the corresponding temperature gradient
Nickel-Chromium Alloy
Copper-Nickel Alloy
The relationship between temperature difference and voltage varies with materials
+
-
VS = SΔT

## Slide 40: Measuring Temperature
To measure temperature using a thermocouple, you can’t just connect the thermocouple to a measurement system (e.g. voltmeter)
SOURCE: http://www.pcbheaven.com/wikipages/images/thermocouples_1271330366.png

## Slide 41: Measuring Temperature
To measure temperature using a thermocouple, you can’t just connect the thermocouple to a measurement system (e.g. voltmeter)
The voltage measured by your system is proportional to the temperature difference between the primary junction (hot junction) and the junction where the voltage is being measured (Ref junction)
SOURCE: http://www.pcbheaven.com/wikipages/images/thermocouples_1271330366.png

## Slide 42: Measuring Temperature
To measure temperature using a thermocouple, you can’t just connect the thermocouple to a measurement system (e.g. voltmeter)
The voltage measured by your system is proportional to the temperature difference between the primary junction (hot junction) and the junction where the voltage is being measured (Ref junction)
SOURCE: http://www.pcbheaven.com/wikipages/images/thermocouples_1271330366.png
To determine the absolute temperature at the hot junction…
You need to know the temperature at the Ref junction!

## Slide 43: Measuring Temperature
To measure temperature using a thermocouple, you can’t just connect the thermocouple to a measurement system (e.g. voltmeter)
The voltage measured by your system is proportional to the temperature difference between the primary junction (hot junction) and the junction where the voltage is being measured (Ref junction)
SOURCE: http://www.pcbheaven.com/wikipages/images/thermocouples_1271330366.png
To determine the absolute temperature at the hot junction…
You need to know the temperature at the Ref junction!
How can we determine the temperature at the reference junction?

## Slide 44: Ice Bath Method (Forcing a Temperature)
Thermocouples measure the voltage difference between two points
To know the absolute temperature at the hot junction, one must know the temperature at the Ref junction

## Slide 45: Ice Bath Method (Forcing a Temperature)
Thermocouples measure the voltage difference between two points
To know the absolute temperature at the hot junction, one must know the temperature at the Ref junction
NIST thermocouple reference tables are generated with Tref = 0 oC
Vmeas = V(Thot) – V(Tref)
V(Vhot) = Vmeas + V(Tref)
If we know the voltage-temperature relationship of our thermocouple, we could determine the temperature at the hot junction
IS IT REALLY THAT EASY?

## Slide 46: Nonlinearity in the Seebeck Coefficient
Thermocouple output voltages are highly nonlinear
The Seebeck coefficient can vary by a factor of 3 or more over the operating temperature range of the thermocouples
VS = SΔT

## Slide 47: Temperature Conversion Equation
T = a0 + a1V + a2V2 + …. + anVn

## Slide 48: Look-Up Table for a Type T Thermocouple
Voltage difference of the hot and cold junctions: VD = 3.409 mV
What is the temperature of the hot junction if the cold junction is at 22 oC?

## Slide 49: Look-Up Table for a Type T Thermocouple
Voltage difference of the hot and cold junctions: VD = 3.409 mV
What is the temperature of the hot junction if the cold junction is at 22 oC?
At 22 oC, the reference junction voltage is 0.870 mV
The hot junction voltage is therefore 3.409 mV + 0.870 mV = 4.279 mV
The temperature at the hot junction is therefore 100 oC

## Slide 50: APPLYING WHAT WE’VE LEARNED
Voltage difference of the hot and cold junctions: VD = 4.472 mV
What is the temperature of the hot junction if the cold junction is at –5 oC?

## Slide 51: APPLYING WHAT WE’VE LEARNED
Voltage difference of the hot and cold junctions: VD = 4.472 mV
What is the temperature of the hot junction if the cold junction is at –5 oC?
At -5 oC, the cold junction voltage is –0.193 mV
The hot junction voltage is therefore 4.472 mV – 0.193 mV = 4.279 mV
The temperature at the hot junction is therefore 100 oC

## Slide 52: Is This Really Practical For a Rocket?
What is another method of determining the temperature at the reference junction?

## Slide 53: Cold Junction Compensation
SOURCE: http://www.industrial-electronics.com/DAQ/images/10_13.jpg

## Slide 54: Cold Junction Compensation
SOURCE: http://www.industrial-electronics.com/DAQ/images/10_13.jpg
How could I determine the temperature of the block?

## Slide 55: Cold Junction Compensation
SOURCE: http://www.industrial-electronics.com/DAQ/images/10_13.jpg

## Slide 56: Acquiring Data

## Slide 57: Temperature Measurement Devices in Lab
>

## Slide 58: Resistive Temperature Detector (RTD)
Two terminal device
Usually made out of platinum
Positive temperature coefficient
Tends to be linear
R = R0(1+α)(T-T0) where T0 = 0oC
R0 = 100 Ω, α = 0.03385 Ω/ Ω oC
At 10oC, R = 100(1+0.385)(10) = 103.85 Ω
They are best operated using a small constant current source
Accuracy of 0.01 oC
EXPENSIVE!
SOURCE: http://www.omega.com/prodinfo/images/RTD_diag1.gif

## Slide 59: Temperature Measurement Devices
>

## Slide 60: How Do I Know If These Are Working?
SOURCE: http://elcodis.com/photos/19/51/195143/to-92-3_standardbody__to-226_straightlead.jpg
SOURCE: http://www.accuglassproducts.com/product.php?productid=17523
SOURCE: http://www.eng.hmc.edu/NewE80/PDFs/VIshayThermDataSheet.pdf

## Slide 61: Calibration
How could we calibrate a temperature sensor?

## Slide 62: Calibration
How could we calibrate a temperature sensor?
25 oC
0 oC
100 oC

## Slide 63: Calibration
How could we calibrate a temperature sensor?
25 oC
0 oC
100 oC
SOURCE: http://www.thermoworks.com/products/calibration/usb_reference.html
USB Reference Thermometer

## Slide 64: Calibration
How could we calibrate a temperature sensor?
25 oC
0 oC
100 oC
Each probe includes an individual NIST-Traceable calibration certificate with test data at 0, 25, 70, and 100°C.
SOURCE: http://www.thermoworks.com/products/calibration/usb_reference.html

## Slide 65: Tracking the Rate of Temperature Change
If a slow sensor is placed into a rocket that is launched to a high altitude, the sensor may not be able to track the rate of temperature change
A critical property of a temperature-measurement device is how quickly it responds to a change in external temperature

## Slide 66: Thermal System Step Response

## Slide 67: Thermal System Step Response

## Slide 68: Thermal System Step Response

## Slide 69: Thermal System Step Response

## Slide 70: Thermal System Step Response
Thermal Time Constant
The thermal time constant can be measured as the time it takes to get to (1/e) of the final temperature
100 (1-(1/e)) = 63 oC

## Slide 71: Thermal System Step Response
Thermal Time Constant
The thermal time constant can be measured as the time it takes to get to (1/e) of the final temperature
100 (1-(1/e)) = 63 oC

## Slide 72: Thermal System Step Response
http://www.eng.hmc.edu/NewE80/PDFs/TemperatureMeasurementLecNotes.pdf
http://www.eng.hmc.edu/NewE80/PDFs/TemperatureMeasurementLecNotes.pdf
http://www.colorado.edu/MCEN/Measlab/background1storder.pdf

## Slide 73: SUMMARY
  Measuring Temperature
  Types of Temperature Sensors
    Thermistor
    Integrated Silicon Linear Sensor
    Thermocouple
    Resistive Temperature Detector (RTD)
  Choosing a Temperature Sensor
  Calibrating Temperature Sensors
  Thermal System Transient Response

## Slide 74: References
Previous E80 Lectures and Lecture Notes
  http://www.eng.hmc.edu/NewE80/TemperatureLec.html
Thermcouples White Paper
  http://www.ohio.edu/people/bayless/seniorlab/thermocouple.pdf (downloaded 02/04/2015)
University of Cambridge Thermoelectric Materials for Thermocouples
  http://www.msm.cam.ac.uk/utc/thermocouple/pages/ThermocouplesOperatingPrinciples.html (viewed 02/04/2015)
National Instruments Temperature Measurements with Thermocouples: How-To Guide
  http://www.technologyreview.com/sites/default/files/legacy/temperature_measurements_with_thermocouples.pdf (downloaded 02/04/2015)
Vishay NTCLE100E3104JB0 Data Sheet
  http://www.eng.hmc.edu/NewE80/PDFs/VIshayThermDataSheet.pdf (downloaded on 02/04/2015)